上海昊量光电设备有限公司新产品添加表格

| 网页填写部分 |
| --- |
| 填表人 | Sophia |
| 产品中文名称 | 780nm高脉冲能量飞秒激光器（400 n J） |
| 产品英文名称 | Calmar_laser_Cazadero_780nm_high_pulse_energy_femtosecond_fiber_laser |
| 一级分类 | 激光器 |
| 二级分类 | 皮秒/飞秒激光器 |
| 所属品牌 | 无品牌（注：如不需要标明品牌的，则填写“无品牌”） |
| 产品预览图 | 注：此处不要放很大的图片。如果图比较大，请把图片像素降下来，大小控制在20k以下） |
| 是否需要上载数据单 | 是，Calmar-780nm高脉冲能量飞秒激光器（400 n J） （注：数据单名称编写规则：公司名-产品名，例如：Nanobase-Xperay tunable laser。如果公司名不方便提供，公司名则用Aunion代替） |
| 产品简介: | 美国CALMAR LASER公司CAZADERO系列780nm飞秒激光器，是一系列基于飞秒光纤啁啾脉冲放大技术（FLCPA）的780nm高脉冲能量飞秒激光器，最高输出能量可达400n J。 美国CALMAR LASER公司CAZADERO系列780nm飞秒激光器的种子光是一款被动锁模光纤激光器。通过脉宽展宽将种子光发出的短脉冲展宽以使峰值功率降低，从而避免了啁啾放大过程（CPA）中对光学器件的损伤，展宽后的种子光输入光纤放大器进行放大。从而实现最高可达400n J的单脉冲能量、小于450fs的脉冲宽度。根据客户的实际需求可通过激光器内部压缩实现小于90 fs的窄脉冲宽度！ (限170字之内，对产品概括性的描述，务必包含主要关键词) |
| 产品广告语： | 基于FLCPA 780nm高脉冲能量(400nJ)飞秒激光器（1560、390nm 波长可选）！最短脉冲宽度可达90fs! （请注意是广告语，所以要简短有力，突出重点，让人印象深刻） |
| 产品别称及关键字： | 飞秒激光器，大功率飞秒激光器，1560nm飞秒激光器，780nm飞秒激光器 ，390nm飞秒激光器，CAZADERO， FLCPA-01C-780，飞秒光纤激光器，高功率飞秒脉冲激光器，超快光纤激光器,飞秒光纤激光器，太赫兹产生 |
| 网页主体内容 | FLCPA-01C-780系列780nm飞秒激光器，是基于飞秒光纤啁啾脉冲放大技术（FLCPA）的780nm高脉冲能量飞秒激光器，最高输出能量可达400n J。 FLCPA-01C-780系列780nm飞秒激光器的种子光是一款被动锁模光纤激光器。通过脉宽展宽将种子光发出的短脉冲展宽以使峰值功率降低，从而避免了啁啾放大过程（CPA）中对光学器件的损伤，展宽后的种子光输入光纤放大器进行放大。从而实现最高可达400n J的单脉冲能量、小于450fs的脉冲宽度。 FLCPA-01C-780系列高脉冲能量780nm飞秒光纤激光器的重频及单脉冲能量均可由用户自行设定，脉冲宽度设定范围为90~450 fs。在不同的脉冲重复频率下，FLCPA-01C-780系列均可实现最优化的脉宽及脉冲形状。这款飞秒激光器提供脉冲选择器（pulse picker）选项，可调轻松控制输出脉冲的频率，您甚至可以拾取单脉冲进行输出。此外我们还提供功率控制选项，以方面客户实时的对功率及脉冲能量进行控制，还可提供burst mode控制模式。 如果您在所选脉冲频率下需要更高的脉冲能量，FLCPA-01C-780系列高脉冲能量780nm飞秒光纤激光器可根据您的需求进行优化！ 如果您需要更短的脉冲输出，FLCPA-01C-780系列高脉冲能量780nm飞秒光纤激光器还可以提供脉冲压缩的选项，在保持单脉冲能量的情况下，脉冲最短可被压缩至90fs！ FLCPA-01C-780系列高脉冲能量780nm飞秒光纤激光器内部包含了1560nm，可手动在1560nm和780nm两种波长之间进行切换！双波长可以从同一个激光口共线输出，也可以从两个不同的输出口输出，极大的方便客户的各种使用要求。 FLCPA-01C-780系列780nm飞秒光纤激光器广泛用于生物领域、激光加工、材料分析、基础研究等领域成为飞秒固体激光器的理想替代产品。 主要特点： 高脉冲能量输出：400 n J 波长可选：1560nm/780nm/390nm 脉宽可调：90~450 fs 脉冲能量可调 重频可调，直至单脉冲输出 终身免维护 全光纤结构，可靠性更高 TEM00模输出， M2 < 1.2 线性偏振输出 脉冲压缩选项 主要应用： 钛宝石激光器种子源 生物光子学 太赫兹产生 材料分析 光学测量 多光子显微 参数如下： Model Number FLCPA-01C-780-2 FLCPA-01C-780-10 FLCPA-01C-780-2C Optical Specifications Output Wavelength (nm) 780 Minimum Pulse Width (fs) <450 <450 <90 Maximum Pulse Energy (n J) >400@2MHz >80@10MHz >280@2MHz Repetition Rates and Pulse Energies 16 repetition rates from 120 KHz to 10 MHz。Pulse energy >400 nJ from 120 KHz to 2MHz. Pulse energy drop by power 2 W from >2-10 MHz。 16 repetition rates from 120 KHz to 10 MHz。Pulse energy >80 nJ from 120 KHz to 10 MHz. 16 repetition rates from 120 KHz to 10 MHz。Pulse energy >280 nJ from 120 KHz to 2MHz. Pulse energy drop by power 2 W from >2-10 MHz。 Maximum Average Power (W) >0.8@2 MHz >0.8@10 MHz >0.56@2MHz Beam Quality, M2 <1.2 Polarization Extinction Ratio (dB) >20 Beam Size at output (mm) 1 (典型） Output Beam Free space output, collimated beam Electrical Specifications Voltage (V) 85-264 VAC and 50/60 Hz frequency Electrical Synchronization Output ~ 0.5 V,SMA connector Power Consumption (VA) <350 (100 typical) Environmental Specifications Cooling Air cooling Operating temperature (℃） 17-32 Storage Temperature (℃） 0-50 Physical Specifications Laser Head Dimensions (cm, Kg) 48(W) x76(D) x14(H); ~31 Kg Electrical Controller (cm) 48(W) x50(D) x18(H); ~14 Kg Connections ~ 2 meter long electrical cables I/O Control Specifications Diagnosis Internal self diagnostics via RS232 Laser Status Indicators on Controller front panel Electrical Power On/Off, Laser On&Good Emergency Stop 可选波长390 nm的参数如下： Model Number FLCPA-01C-390-2 FLCPA-01C-390-10 FLCPA-01C-390-2C Optical Specifications Output Wavelength (nm) 390 Minimum Pulse Width (fs) <400 （典型） <400 <90 Maximum Pulse Energy (n J) >150@2MHz >30@10MHz >90@2MHz Repetition Rates and Pulse Energies 16 repetition rates from 120 KHz to 10 MHz。Pulse energy >150 nJ from 120 KHz to 2MHz. Pulse energy drop to 0 from >2-10 MHz。 16 repetition rates from 120 KHz to 10 MHz。Pulse energy >30 nJ from 120 KHz to 10 MHz. 16 repetition rates from 120 KHz to 10MHz。Pulse energy >150 nJ from 120 KHz to 2MHz. Pulse energy drop to 0 from >2-10 MHz。 Maximum Average Power (W) >0.3@2 MHz >0.3@10 MHz >0.18@2MHz Beam Quality, M2 <1.3 (典型） Polarization Extinction Ratio (dB) >20 Beam Size at output (mm) 1 (典型） Output Beam Free space output, collimated beam Electrical Specifications Voltage (V) 85-264 VAC and 50/60 Hz frequency Electrical Synchronization Output ~ 0.5 V, SMA connector Power Consumption (VA) <350 (100 typical) Environmental Specifications Cooling Air cooling Operating temperature (℃） 17-32 Storage Temperature (℃） 0-50 Physical Specifications Laser Head Dimensions (cm, Kg) 48(W) x76(D) x14(H); ~31 Kg Electrical Controller (cm) 48(W) x50(D) x18(H); ~14 Kg Connections ~ 2 meter long electrical cables I/O Control Specifications Diagnosis Internal self diagnostics via RS232 Laser Status Indicators on Controller front panel Electrical Power On/Off, Laser On&Good Emergency Stop 可选波长1560 nm的参数如下： Model Number FLCPA-01C-1560-2 FLCPA-01C-1560-10 FLCPA-01C-1560-2C Optical Specifications Output Wavelength (nm) 1560 Minimum Pulse Width (fs) <500 （典型375） <500 （典型375） <90 Maximum Pulse Energy (n J) >1000@2MHz >200@10MHz >700@2MHz Repetition Rates and Pulse Energies 16 repetition rates from 120 KHz to 10 MHz。Pulse energy >1 uJ from 120 KHz to 2MHz. Pulse energy limited by 2 W from >2-10 MHz 16 repetition rates from 120 KHz to 10 MHz。Pulse energy >200 nJ from 120 KHz to 10 MHz 16 repetition rates from 120 KHz to 10MHz。Pulse energy >700 nJ from 120 KHz to 2MHz. Pulse energy limited by 2 W from >2-10 MHz Maximum Average Power (W) >2 >2 >1.4 Beam Quality, M2 <1.2 Polarization Extinction Ratio (dB) >20 Beam Size at output (mm) 3 (典型） Output Beam Free space output, collimated beam Electrical Specifications Voltage (V) 85-264 VAC and 50/60 Hz frequency Electrical Synchronization Output ~ 0.5 V, SMA connector Power Consumption (VA) <350 (100 typical) Environmental Specifications Cooling Air cooling Operating temperature (℃） 17-32 Storage Temperature (℃） 0-50 Physical Specifications Laser Head Dimensions (cm, Kg) 48(W) x76(D) x14(H); ~31 Kg Electrical Controller (cm) 48(W) x50(D) x18(H); ~14 Kg Connections ~ 2 meter long electrical cables I/O Control Specifications Diagnosis Internal self diagnostics via RS232 Laser Status Indicators on Controller front panel Electrical Power On/Off, Laser On&Good Emergency Stop （产品内容，添加时注意图片转成JPG，尺寸不能太大，太大的图片缩小在放上去） |
| 需要添加的相关产品 | http://www.auniontech.com/p/ultrafast_laser/Calmar_Carmel_780nm_1550nm_fs_laser.html http://www.auniontech.com/p/ultrafast_laser/Calmar_laser_Mendocino_780nm_picosecond_fiber_coupled_laser.html http://www.auniontech.com/p/ultrafast_laser/Calmar_Laser_Mendocino_780nm_femtosecond_Fiber_Laser_With_Free_Space_Output.html http://www.auniontech.com/p/ultrafast_laser/Aligna_Automated_Laser_Beam_Alignment_and_Stabilization_System.html http://www.auniontech.com/p/ultrafast_laser/ConOptics_Modulation_System_for_Multiphoton_Microscopy.html （除非同批次还没添加的新产品，网站上已添加的的老产品请务必列出其链接） |
| 是否需要添加FAQ |  |
| 需要链接此产品的产品 | http://www.auniontech.com/p/ultrafast_laser/Calmar_laser_Cazadero_1550_high_energy_femtosecond_laser.html（http://www.auniontech.com/p/ultrafast_laser/Calmar_Carmel_780nm_1550nm_fs_laser.html （去掉Conoptics 双光子显微钛宝石激光器专用分束器） （如果五个相关产品已满，请指出希望替代那一项） |

| 百度推广部分 |
| --- |
| 是否要进行搜索引擎推广 | 是 | 备注： |
| 搜索关键词 | 飞秒激光器，大功率飞秒激光器，1560nm飞秒激光器，780nm飞秒激光器 ，390nm飞秒激光器，CAZADERO， FLCPA-01C-780，飞秒光纤激光器，高功率飞秒脉冲激光器，超快光纤激光器,飞秒光纤激光器，太赫兹产生 | 请填写客户可能搜索的所有关键词）按倒叙和重要程度写 |
| 搜索结果标题： |  | （20个汉字之内） |
| 产品描述第一行： | 美国CALMAR LASER公司CAZADERO FLCPA-01C-780系列飞秒激光器，是基于飞秒光纤啁啾脉冲放大技术（FLCPA）的780nm高脉冲能量飞秒激光器，最高输出能量可达400n J。 | （40个汉字之内，需要是完整而独立的句子，不允许出现依托于第二句才能表达完整的意思的情况）里面最好要包含最重要的关键字 |
| 产品描述第二行： | 美国CALMAR LASER公司FLCPA-01C-780系列高脉冲能量飞秒光纤激光器内部包含了1560nm、390nm波长可手动切换选择输出。 | （40个汉字之内）里面最好要包含最重要的关键字 |