EMS

EMS 850 Critical Point Drier

Instruction Manual

Electron Microscopy Sciences

1560 Industry Road,

Hatfield, Pa 19440

USA

Tel          +1-215-412-8400

Fax         +1-215-412-8450

E-Mail     sgkcck@aol.com

INDEX

1.  DESCRIPTION	4

2.  INSTALLATION	5

2.1	Preliminary Checks	5

2.2	Connections.	6

2.3	Controls	7

2.4	Initial Operating Checks.	8

3.  OPERATION	10

3.1	Specimen Preparation	10

3.1.1 Fixation	10

3.1.2 Dehydration	10

3.1.3 Intermediate	10

3.1.4 Transitional fluid	10

3.2	Specimen Handling	10

3.3	Critical Point Drying	11

3.4	Safety Devices.	12

3.5	Safety Precautions	12

4.  SUMMARY	13

5.  SERVICE AND MAINTENANCE	14

5.1	Maintenance	14

6.  SPARES AND ACCESSORIES	15

7.  Appendix	16

7.1	Figure 1, Rear Panel	16

7.2	Figure 2, Front Panel	17

7.3	Figure 3, Cross-section View Of Chamber	18

7.4	Figure 4, Chamber Fill Level	19

7.5	Fuse Listing  - EMS850	20

7.6	Electrical Supplies Worldwide.	21

7.7	Cylinder valve outlet & Hose connections	22

8.  REFERENCES	23

9.  DOCUMENT HISTORY	24

C E Declaration:

This Equipment of this Design and manufacture and marked CE, conforms with the requirements of the European Directives EMC 89/336/EEC & LVD 73/23/EEC.

This Equipment will “fail safe” in the presence of excessive RF, Electrostatic Discharge or Mains Transients.  While a loss of function could occur under extreme circumstances the Equipment’s operation will be fully recoverable under normal operating conditions.

Labelling.

Mains Lead.

This Equipment must be Earthed and fitted with the correct lead for the country of operation.  This will normally be achieved from the correct mains supply socket

Earth Connector.

This Equipment is normally supplied from 3-pin supply including Earth.  If only 2-pin supply is available a separate Earth must be fitted.  The supplementary Earth stud can be used to facilitate this requirement.

Inlet/Exhausts:

The applications of these 4 gas connections are given in the layout Appendix 7.1, Figure 1.  The gas supply from high-pressure cylinder must be fitted with the correct fitting for the cylinder type.  The exhaust, vent & bleed - should be safely vented.

# 1.  DESCRIPTION

Critical point drying is so named as it includes, as part of its process, the occurrence known as the continuity of state for which there is no apparent difference between the liquid and gas state of a medium, the surface tension between this interface reducing to zero.  This occurs at a specific temperature and pressure with resulting density and is known as the Critical Point.  This condition of zero surface tension can be used to dry biological specimens, avoiding the damaging effects of surface tension.

In biological specimens we are mainly concerned with the removal of water.  Unfortunately, the critical point for water (H2O) of +374°C and 3212psi is inconvenient, and would cause heat damage to the specimen.  The most common and convenient transitional medium for Critical Point Drying is Carbon Dioxide (CO), which has its Critical Point at 31°C and 1072psi.  However, it is not miscible with water, and therefore we have to involve a third medium, commonly Acetone, which is termed the intermediate fluid.  We can now convert our transitional fluid, typically CO2, from liquid to gas without surface tension at the Critical Point.

The EMS850 Critical Point Drier is designed for use with CO2 having first replaced any water in the specimen by a series of dehydration, often in the same fluid such as Acetone or Ethanol, which will also be the intermediate fluid.

(Wet Specimen) - water - Acetone - 30% - 100% - CO2 - C.P.D. - (Dry Specimen)

The specimens for Critical Point Drying are located in the pressure chamber of the EMS850.  The chamber is pre-cooled to allow it to be readily filled with liquid CO2 from a gas cylinder.  The chamber is then heated to just above the Critical Temperature with subsequent Critical pressure being achieved.  The CO2 gas is vented through a needle valve, to avoid specimen distortion.

The EMS850 is fitted with thermoelectric heating and adiabatic cooling, and temperature control of +35°C on heating with associated pressure on heating of the order of 1250psi.  This ensures the Critical Point is accurately obtained, avoiding excess pressures or temperatures, or the need to rely on pressure relief valves to control pressure during the heating cycle.  The chamber is vertical, with top loading, to ensure specimens do not become uncovered during the drying process, and has a side viewing port to locate the meniscus for the correct level when initially filling the chamber.

# 2.  INSTALLATION

It is important that this Equipment is installed and operated by skilled personnel in accordance with these Instructions.  Failure to do so may result in damage, and impair protection provided.

'If in doubt - ASK'.

A suitable location should be provided for the unit - either operated on a bench or the recommended trolley.  The total weight of the system is 18Kg.  The system operating environment ambient temperature range is 15oC to 25oC in a non-condensing relative humidity of not more than 75%.  Sufficient ventilation is required, and positioning should be out of direct sunlight.  The system is rated for continuous operation other than those supplies specified.

## 2.1	Preliminary Checks

Remove Instrument from packing and place on appropriate operational position.  Carry out visual inspection for any signs of transit damage.  Particular attention should be given to inspecting the high pressure connecting hose for signs of damage.

In choosing an operational position it should be noted that for normal operation the pre-cooling and heating parameters might be affected by changes in ambient temperature.  The Instrument should not be located near sources of heat or cold, or any other affects such as direct sunlight.  The area should also be of relatively low humidity to avoid moisture condensation within the pre-cooled chamber when loading specimens.  Also allow for safe venting of gas.  The gas cylinder of Carbon Dioxide should be located adjacent to the right hand side of the Instrument, and should be securely located.

Remove Accessory Pack, and check contents against EMS850 Despatch Certificate.  Ensure that the Independent Test Certificate is with the Instrument as proof of pressure testing.

Ensure that all areas of the instrument are free of loose packaging material.

Check that the Instrument is of the correct voltage for the Country of operation by checking the serial number plate and the voltage label on the back of the Instrument.

The Instrument is designed for use with Acetone or Ethanol as the intermediate fluid, and Carbon Dioxide (via high pressure gas cylinder) as the transitional fluid.  Research grade Carbon Dioxide gas should be used and the cylinder fitted with a siphon tube (usually denoted by a White stripe on a Black cylinder).

## 2.2	Connections.

WARNING: THIS EQUIPMENT CAN PRODUCE HIGH VOLTAGES AND HIGH PRESSURES

IN ITS OPERATION.

Connections should only be made in accordance with Instructions.

UNDER NO CIRCUMSTANCES SHOULD ANY OTHER CONNECTIONS OR OUTLETS/INLETS BE USED FOR ANY OTHER EQUIPMENT OR SERVICES.

Refer to:  Appendix 7.1, Figure 1: Rear Panel Diagram.

| Title | Function |
| --- | --- |
| Rocker switch | Main power on to Instrument. |
| Neon | Main power on indication |
| Power in | Main power inlet socket |
| Gas supply | High pressure gas inlet (carbon dioxide) |
| Exhaust | Gas outlet direct from chamber. |
| Bleed | Gas outlet via flow gauge from chamber. |
| Vent | Gas outlet from Adiabatic cooling system. |
| Voltage | Correct voltage for Country of operation. |
| Fuse 1 | Main power fuse |

For use ratings and voltages refer to Appendix 7.2.

Carry out high-pressure connections to the Instrument with the flexible braid covered high-pressure hose provided.

The Instrument connection is common, and is a direct metal contact via a small nipple located in the gas supply inlet.  Tighten up firmly, using the correct size spanner which, used normally, will ensure the correct torque is applied.  The gas cylinder connection will vary depending on the Country of operation, and the type of gas cylinder used.

CAUTION: Ensure you have correct connection - if in any doubt -do not fit.  (Connector types are given in Appendix 7.4)

The gas cylinder connection uses a fibre washer, which is provided and should be placed on the end of the flat face of the connector, inside the nut housing, and matched with the flat surface of the connection on the high-pressure gas cylinder.  The nut should be tightened up using the correct size spanner which, used normally, will ensure the correct torque is applied.

CAUTION:  Ensure that the gas cylinder is secured and the valve turned off.

The electrical input to the Instrument is made with the power lead provided.  The Instrument connection is standard and the lead is fitted with the appropriate plug for the Country of operation.

(Electrical supplies are given in Appendix 7.3)

NOTE:	A single phase AC supply with Earth is required - selected to the correct voltage for the Country of operation, either nominal 240V or nominal 120V.  The voltage and frequency range is:

| Nominal 240 | Max current 4A | 200 - 264V | 47Hz to 63Hz |
| --- | --- | --- | --- |
| Nominal 120 | Max current 8A | 90v - 132V | 47Hz to 63Hz |

Any outlet gases from the Instrument should be safely vented.  The small-bore tubing provided can be utilized.

The vent and bleed can be connected using the 'T' piece provided.  The exhaust has separate tubing; this is to allow it to be removed more easily when checking of intermediate fluid purging.  (Described in a later section).

If the optional flow meter is required this should be connected to the bleed outlet using the small bore tubing provided.  Care should be taken not to trap any of the tubing, which should be secured at intervals along its length.

## 2.3	Controls

Refer to:  Appendix 7.1, Figure 2:  Front Panel Diagram.

| Title | Function |
| --- | --- |
| Cool Valve (Blue) | Allows controlled flow of gas to adiabatic cooling system - the change in pressure giving a cooling effect to the chamber - the adiabatic cooling system does not let gas into the chamber. |
| Inlet valve (Green) | Allows gas direct to chamber |
| Exhaust Valve (Black) | Allows direct exhausting of gas from chamber |
| Bleed Valve (Red) | Allows bleeding of gas from chamber via a flow meter for controlled letdown |
| Temperature Gauge | Monitors temperature of chamber |
| Pressure Gauge ** | Monitors direct pressure in chamber |
| Stirrer Switch | Switches on motor, which is located on the base of chamber, this has a magnetic coupling to a stirrer bar located inside the chamber. |
| 12V DC LED | Indicates 12V dc control voltages are present. |
| Heater Control LED | Indicates when heater is on or off under control mode |
| Heater Switch | Switches power on to heater control system |
| Flow Gauge (Optional) | Allows monitor and fine control of letdown |

**NOTE:  This type of pressure gauge is filled with a 'damping' fluid, often evident by a small bubble in the gauge dial.

## 2.4	Initial Operating Checks.

(These should be made having become familiar with the controls.)

Ensure that the high-pressure gas cylinder valve is closed.  Check that all valves can be opened and closed freely.  Close inlet valve and open exhaust valve.  Remove chamber lid by unscrewing the three thumbnuts.

CAUTION:	When the chamber lid is to be removed or refitted - ensure inlet valve is closed, exhaust valve is open, and pressure gauge is reading zero.

Position the stirrer bar in the bottom of the chamber.  Switch on stirrer, and observe rotation.

Refit the chamber lid, ensuring the 'O' ring is not damaged and is correctly seated.  Tighten thumbnuts firmly in even rotational sequence.

NOTE:	The front viewing window assembly includes a quartz glass window and a polycarbonate (Lexan) shield in front of it for additional safety protection.  Removing or refitting should be carried out with the correct replacement parts, and in accordance with separate detailed Instructions.

Switch on heater.  The heater control should illuminate, and the temperature will increase gradually at approximately 2°C/minute.  The chamber will feel slightly warm.  Switch off heater when initial indications are that it is operating.  Close all valves, and open gas cylinder valve.  Check for any obvious signs of leaks.

Open cool valve by a very small amount until a steady hissing of gas is established.  After a short delay, the temperature will start to fall.  The cool down rate is approximately 4oC/minute.  Shut cool valve when indications are that it is operating.

Open inlet valve.  The pressure gauge should read approximately 800psi.  Close the inlet valve, and depressurise the chamber firstly by using the bleed valve and checking the flow gauge if fitted is operational, and then by using the exhaust valve for quick depressurising.  This will also ensure that the system has been 'flushed through'.  Close the bleed valve and the exhaust valve.

TEST RUN

Open cool valve slowly to establish a steady flow.

The chamber, after an initial thermal delay, should cool down from an ambient temperature of +20°C to +5°C in approx. 3minutes.  At +5°C the cool valve can be closed and the temperature will 'coast' down to approx. +2°C.  With practice, a cool valve setting can be achieved which avoids freezing of the valve and internal pipe-work, thus avoiding the  'spitting' associated with occasional freezing.

Open the inlet valve slowly, taking approx. 1minute to fill the chamber to the correct level by observing the meniscus in the viewing window.  On filling, the temperature may increase to +5°C.  If it gets any warmer, and any difficulty is experienced with filling, then a small amount of cooling can be applied.  The chamber should be filled to the centre of the viewing window to filling mark shown.

IMPORTANT	DO NOT over fill the chamber.  This causes excessive pressures, and whilst the safety pressure relief valve will operate, it is disruptive for the specimens.  (See Appendix 7.1 Fig.4)

Close the inlet valve and switch on the heater.

NOTE: The high-pressure carbon dioxide cylinder valve can also be closed for additional safety.

After an initial delay the temperature will increase to the control temperature of +35°C, this takes approximately 15 minutes.  At this point it will control switching in and out over the range 32°C TO 38°C and pressure in the range of 1200 to 1300psi.  If the meniscus is observed, it will fall slightly to about half way in the viewing window.

At a temperature of approximately 30°C, and a corresponding pressure of 1,100psi the meniscus will disappear having reached the Critical Point.

NOTE:	Subsequent cooling would allow the re-appearance of the meniscus, should this be of interest for demonstrations.

The system can now be depressurised whilst maintaining the temperature.

For less delicate specimens the exhaust valve can be used at a rate of 100psi/minute, taking approximately 10 minutes.

For more delicate specimens the bleed valve can be used, for setting a lower de-pressuring rate of 1000cm3/Minute (2 litre/minute) taking approximately 20 minutes.  (An optional flow meter can be connected to the bleed outlet to assist in setting the de-pressuring rate.)

The heater can now be switched off.

If a run has to be repeated immediately, then the cooling from the operational temperature of +35oC will take somewhat longer.

If the instrument has to be shut down - the carbon dioxide cylinder valve should be closed. The cool and exhaust valves open to ensure depressurisation of the lines and the chamber.

Mains power switched off.

# 3.  OPERATION

## 3.1	Specimen Preparation

A range of techniques may be used for specimen preparation and the following in only intended as a general guide.

### 3.1.1 Fixation

Fixation generally utilizes an Aldehyde, such as Formaldehyde or Glutaraldehyde to provide a rapid inter and intra-cellular penetration and fixation, followed by Osmium Tetroxide for further fixation.

### 3.1.2 Dehydration

The fixed specimen is taken through a series of increasing concentrations of a drying liquid such as Acetone or Ethanol, ending in a 100% dehydrating liquid of the highest possible purity.  Typically these are steps of 10, 20, 30, 50, 70, 90, 100% at 10 minutes each, with 3 changes at 100%.

### 3.1.3 Intermediate

The intermediate fluid, which is miscible with Carbon Dioxide, is typically the dehydration fluid and therefore an additional step to the intermediate fluid, prior to Critical Point Drying, is not required.  The most common are Acetone and Ethanol.

### 3.1.4 Transitional fluid

Carbon Dioxide - Critical Temperature 31.1°C, Critical pressure 1072psi.

## 3.2	Specimen Handling

Generally the wire mesh specimen basket is most commonly used when transferring the smaller specimen.  For larger specimens, the porous pots may be used.  The specimen must be maintained in the dehydration fluid immediately prior to use, then quickly transferred to the EMS850 chamber.

NOTE:	To avoid specimens becoming uncovered when transferred to the EMS850 chamber, they should only occupy two thirds of the chamber.  (See Appendix 7.1 Fig.3)

## 3.3	Critical Point Drying

NOTE: Ensure all valves are closed.

- Pre cool the chamber to +5°C using the cool valve as previously described.

- Transfer the specimen, avoiding any unnecessary delay.

- The inlet valve is opened and the CO2 level is filled to the top of sight glass, after soaking, fresh CO2 can be added by opening the exhaust draining the unit to the bottom or even below the sight glass.  This is repeated until the solvent exchange is completed.  The Chamber should now be filled for the final time for Critical Point Drying; it is at this stage that it is IMPORTANT to just fill the CO2 to the approx. centre of the sight glass.  This is to avoid excess pressure when heaters are applied.

NOTE:	If pressure is let down too quickly, spurting noises may be heard, this is the freezing up of the exhaust lines, and a slower let down should be used.

- Allow the specimen to soak for approximately 5 minutes (for very large specimens the soak period should be double) during this period the stirrer may be used to enhance solvent exchange.

- The chamber should now be purged for approximately 1 minute.  This is achieved by slowly opening the exhaust valve and observing the meniscus to fall slightly. The inlet valve should now be opened slightly to compensate and balance the system thereby replenishing the chamber with 'fresh' Carbon Dioxide.

NOTE:	The rotating stirrer should be switched off for the purging, as both Acetone and Ethanol are more dense so will go to the bottom of the chamber which is designed to be top filling bottom draining.

- Ensure the meniscus is at the centre of the viewing window.  Repeat the soak as (d).

- A further purging should now be carried out as in (e) during the second purging of 1 minute, a check can be made to see that solvent exchange has been achieved: Remove the plastic pipe from the exhaust outlet.  Using a small disc of filter paper, place into the stream of gas.  The presence of a small damp patch will indicate if solvent is present.  If necessary, repeat a soak and purge step.

- Ensuring meniscus is at the centre of the viewing window, close all valves.  For added safety also close the carbon dioxide cylinder valve.

- Switch on heater and allow stable conditions to be reached of approx. 35°C and 1250psi.  This will take about 15 minutes.

- The chamber may now be depressurised as described previously.

  - Exhaust Valve at 100psi/minute.

  - Bleed valve and needle valve (delicate specimen) at1000°Cm3/minute flow rate.

- Specimen can now be removed and subsequently treated.  In all cases, this should be maintained in dry conditions.

- The EMS850 Instrument can now be re-used or shut down, as previously described.

## 3.4	Safety Devices.

| Safety Devices | Pressure |
| --- | --- |
| Chamber normal maximum working pressure | 1500psi |
| Complete unit workshop tested. | 2250psi |
| Quartz viewing window fitted with polycarbonate shield (Lexan) tested as part of chamber assembly. NOTE: Chamber and Quartz window have on a type test been independently tested to 4370psi. | To 2250psi |
| Temperature controller factory set. | Mid point: 35oC range: 32°C to 38°C. Typical pressure: 1200psi to 1300psi |
| Temperature thermal trip non adjustable | Set at: 40°C, resets at: 24°C |
| Pressure relief valve factory set | 1700psi |
| Rupture disc non adjustable | Burst at: 1900psi |

## 3.5	Safety Precautions

- Locate in suitable area in firm position.

- Ensure correct venting and ventilation.

- Secure gas cylinder.

- Avoid trapping or kinking of high-pressure hose and vent tubing.

- Close cylinder valve when not in use.

- Open all instrument valves when not in use-switch off at mains.

- Ensure all fittings are correct on initial installation.  IF IN ANY DOUBT - PLEASE ASK.

- Wear safety spectacles for additional safety when viewing window direct.

- Advise with notices all staff of presence and nature of equipment.

WARNING - HIGH PRESSURE EQUIPMENT.

- Use only correct replacement parts from manufacturer.  If quartz window or chamber lid assembly shows any sign of damage even if instrument appears to be functioning normally DO NOT USE.

- Use only Acetone or Ethanol as intermediate fluid.

- Use only Carbon Dioxide as transitional fluid.  (Research grade with siphon tube.)

# 4.  SUMMARY

| Procedure | Temperature/Time/Pressure |
| --- | --- |
| 1 | Dehydrate specimens Acetone or Ethanol |  |
| 2 | Ensure all valves are closed. |  |
| 3 | Pre-cool chamber using cool valve (Blue). Avoid excessive flow and freezing up. | +5°C - 4 Minutes |
| 4 | Load specimens ensure thumb-screws correctly tightened. Check chamber is depressurised. |  |
| 5 | Fill chamber, with meniscus to centre of view window using inlet valve (Green) all other valves closed. | 1 Minute |
| 6 | Soak (with stirrer if required). | +5°C - 3 Minutes |
| 7 | Purge chamber, maintaining meniscus at least mid way on view window using slow, steady flow using balance of inlet valve (Green) and exhaust valve (Black). | 1 Minute |
| 8 | Soak (with stirrer if required). If necessary - pre-cool chamber using cool valve (Blue). All other valves closed. | +5°C - 3 Minutes |
| 9 | (Repeat 7). Checking with filter paper at outlet of exhaust at rear of instrument for damp patch indicating solvent exchange condition. | 1 Minute |
| 10 | Fill Chamber to centre of chamber to filling mark shown. Close all valves. Switch on heater. Allow stable conditions. | +35°C - 15 Minutes 1250psi. |
| 11 | De-pressurise system: (a) Exhaust valve (Black) - Non delicate or; (b) Bleed valve (Red) - Delicate | 100psi/Minute +35°C 10 Minutes 1000cm3/Minute +35oC 20 Minutes |
| 12 | Repeat run. Pre cooling. | 8 Minutes +35oC TO 5oC |
| 13 | Shut down instrument. (a) Close gas cylinder valve. (b) Open exhaust valve (Black). (c) Open cool valve (Blue). (d) Switch off mains electricity. |  |

# 5.  SERVICE AND MAINTENANCE

## 5.1	Maintenance

| Procedure | Weekly/Monthly |
| --- | --- |
| Check chamber lid 'O' ring | Weekly |
| Check chamber studs and thread condition | Monthly |
| Check chamber thumb screws and thread condition | Monthly |
| Check condition of viewing window Check tightness of viewing window screws. (Only dismantle and change in accordance with Instructions supplied with replacement parts) | Weekly |
| Check high-pressure hose for signs of damage. | Weekly and/or prior to sustained period of use. |
| Check high-pressure hose connections at instrument and cylinder for tightness. | Weekly and/or prior to sustained period of use. |
| Check valve-locking nuts for tightness (Only dismantle and change in accordance with Instructions supplied with replacement parts). | Weekly |
| Check condition of mains power cord. | Weekly |

CAUTION:	During any maintenance and service activities, ensure mains electrical power is off, and high pressure hose is disconnected.

NOTE:	Consumables and spare items can be obtained from EMS or approved distributor.  Only EMS recommended items should be used. For technical assistance and advice contact EMS

Electron Microscopy Sciences,

1560 Industry Road

Hatfield, Pa 19440

Tel:  +1-215-412-8400

Fax: +1-215-412-8450

** If approved distributor not known - please contact EMS direct for details.

# 6.  SPARES AND ACCESSORIES

| Spares for K850 Critical Point Drier | Quantity |
| --- | --- |
| Chamber Lid 'O' Rings | x5 |
| Knurled Thumb Screws | x3 |
| High Pressure Hose | x1 |
| High Pressure Hose Fibre Washer | x5 |
| Quartz Window ** | x1 |
| Polycarbonate Protective Shield ** | x1 |
| Viewing Window 'O' Ring ** | x5 |
| Viewing Window Washer ** | x5 |
| Needle Valve | x1 |
| Flow Gauge Tube | x1 |
| Rupture Disc Assembly | x1 |

USEFUL ACCESSORIES FOR EMS 850 CRITICAL POINT DRIER

| Accessories | Quantity |
| --- | --- |
| Porous Specimen Pots | x10 |
| Stirrer Bars | x3 |
| Safety Spectacles | x1 |

**

NOTE: In the unlikely event that the quartz window cracks or shows any signs of damage - please inform EMS immediately, and ensure the Instrument is clearly marked: DO NOT USE.

If the polycarbonate protective shield becomes scratched or damaged, it can be replaced without having to replace the quartz window.

These items should only be changed in accordance with Instructions.

# 7.  Appendix

## 7.1	Figure 1, Rear Panel

## 7.2	Figure 2, Front Panel

## 7.3	Figure 3, Cross-section View Of Chamber

## 7.4	Figure 4, Chamber Fill Level

## 7.5	Fuse Listing  -  EMS 850

| Title | Rating | Function |
| --- | --- | --- |
| 240 Volts Fuse 1 | T630mA Ceramic | Mains Power |
| 120 Volts Fuse 1 | T1A Ceramic | Mains Power |

Replacement fuses can be supplied by EMS, or the Approved Distributor**

Electron Micrsocpy Sciences

1560 Industry Road

Hatfield, Pa 19440

USA

TEL:  +1-215-4128400

FAX:  +1-215-412-8450

**If approved Distributor not known - please contact EMS direct for details.

## 7.6	Electrical Supplies Worldwide.

| Country | AC Volt/Freq. |
| --- | --- |
| Brazil | 220V/60Hz |
| Canada | 115V60Hz |
| France | 220V/50Hz |
| Germany | 220V/50Hz |
| India | 230V/50Hz |
| Ireland | 230V/50Hz |
| Israel | 230V/50Hz |
| Italy | 220V/50Hz |
| Japan | 100-200V/50Hz-60Hz |
| Korea | 115V/60Hz |
| Netherlands | 220V/50Hz |
| Norway | 220V /50Hz |
| Pakistan | 230V/50Hz |
| Portugal | 220V/50Hz |
| Scandinavia | 220V/50Hz |
| Singapore | 220V/50Hz |
| Spain | 220V/50Hz |
| Taiwan | 220V/60Hz |
| Turkey | 220V/50Hz |
| United Kingdom | 220-240V/50Hz |
| United States of America | 115V/60Hz |

## 7.7	Cylinder valve outlet & Hose connections

BS No. 8 		UK

France

Germany

Italy

Sweden

CGA 320		Japan

USA

Caution 	It is important to check the Correct Hose Connection to High pressure Gas Cylinder.

If in doubt - Do not fit - Please ask

# 8.  REFERENCES

1.	COHEN, A.L.

"Critical Point Drying - Principles And Procedures."

Electron Microscopy 1979 II 303-324.

(Available on Request)

# 9.  DOCUMENT HISTORY

| Revision | Date | By: | Reason |
| --- | --- | --- | --- |
| Issue 1 | 23/08/02 | PRM | Initial Issue |
| Issue 2 | 29/12/03 | JB | Revision |
| Issue 3 | 06/09/04 | PRM | BS No 8 Image corrected |

Safe handling and storage of gases in cylinders

| Introduction | The following general practices are recommended for the safe handling and storage of high-pressure gaseous and liquefied-compressed gases in transportable containers. Additional precautions may be necessary depending upon the category to which the gas belongs (corrosive, toxic, flammable, pyrophoric, oxidant, radioactive or inert), the individual properties of the gas, and the process in which it is used. |
| --- | --- |
| General | Only trained persons should handle compressed gases. Observe all regulations and local requirements regarding the storage of containers. Do not remove or deface labels provided by the supplier for identification of the container contents. Ascertain the identity of the gas before using it. Know and understand the properties and hazards associated with each gas before using it. Establish and implement plans to cover any emergency situations that might arise. When doubt exists as to the correct handling procedure for a particular gas contact the supplier. |
| Handling and Use | Wear stout gloves. Never lift a container by the cap or guard unless the supplier states it is designed for that purpose. Use a trolley or other suitable device or technique for transporting heavy containers, even for a short distance. Leave valve protection caps (where provided) in place until the container has been secured against a wall or bench or placed in a stand and is ready for use. Never remove valve guards or cages, which permit access to the container valve. This type of valve protection is intended to be permanently attached. Where necessary wear suitable eye and face protection. The choice between safety glasses, chemical goggles, or full-face shield will depend on the pressure and nature of the gas being used. Where necessary for toxic gases see that self-contained positive pressure breathing apparatus or a full-face airline respirator is available in the vicinity of the working area. Check for gas leaks using a suitable method. Flammable and toxic gas monitors are available. Where a gas is toxic check special procedures specified by supplier. Ascertain that an adequate supply of water is available for first aid, fire fighting, or dilution of corrosive material in the event of leakage. Employ suitable pressure regulating devices on all containers when the gas is being emitted to systems with a lower pressure rating than that of the container. Before connecting the container for use ensure that back feed from the system into the container is prevented. Ascertain that all electrical systems in the area are suitable for service with each gas. Never use direct flame or electrical heating devices to raise the pressure of a container. Containers should not be subjected to temperatures above 45oC. Never re-compress a gas or a gas mixture from a container without consulting the supplier. Never attempt to transfer gases from one container to another. Do not attempt to increase liquid draw-off rate by pressurising the container without first checking with the supplier. |

Safety information for the return of Preparation Equipment and Accessories.

General Introduction

The employer (user) is responsible for the health and safety of his employees.  This also applies to all those persons who come into contact with the Preparation Equipment and Accessories either at the user or manufacturer’s premises during repair of service.  The contamination of Preparation Equipment and Accessories has to be declared and the Health and Safety Declaration form completed.

Health and Safety Declaration

Those persons carrying out repair or service have to be informed of the condition of the components.  This is the purpose of the ‘Declaration of Contamination of Preparation Equipment and Accessories.’.

Despatch

When returning equipment the procedures set out in the Operating Instructions must be followed.  For example:

- 	Drain the vacuum pumps.

-	Neutralise the flushing with gas.

-	Remove filter elements.

-	Seal all outlets.

-	Pack glass components safely.

-	Pack loose attachments securely for example stages.

-	Seal in heavy-duty polythene or a bag,

-	Despatch in suitable transport container.

Return Address:

F.A.O.:  The Service Manager,

EMS

1560 Industry Road

Hatfield  Pa 19440

Declaration of Contamination of Preparation Equipment and Accessories.

The repair and/or service of Preparation Equipment and Accessories can only be carried out if a correctly completed declaration has been submitted.  Non-completion will result in delay.  The manufacturer reserves the right to refuse acceptance of consignments submitted for repair or maintenance work where the declaration has been omitted.

This declaration may only be completed and signed by authorised and qualified staff.

| 1. Description of component - Equipment type/model: ____________________________ - Code No:. ___________________________________ - Serial No.: ___________________________________ - Invoice No. (if known) ___________________________ - Delivery Date.: (if known)___________________________ | 2. Reason for return: __________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________________ |
| --- | --- |
| 3. Equipment condition - Has the equipment been used? Yes/No - What type of operating medium was used? _______________________________________________ - Is the equipment free from potentially harmful substances? Yes/No (If Yes go to Section 5) (If No go to Section 4) | 4. Process related contamination of Equipment/ Accessories. - Toxic Yes/No - Corrosive Yes/No - Explosive* Yes/No - Microbiological* Yes/No - Radioactive* Yes/No - Other harmful substances Yes/No |

* We will not accept any Equipment/Accessories, which have been radioactively, explosively, or microbiologically contaminated without written evidence that such Equipment/Accessories have been decontaminated in the prescribed manner.

Please list all harmful substances, gases and dangerous by-products, which have come into contact with the Preparation Equipment and Accessories.

| Trade name Product name Manufacturer | Chemical name and symbol | Danger class | Precautions associated with substance. | First aid measures in the event of an accident. |
| --- | --- | --- | --- | --- |
| 1. |  |  |  |  |
| 2. |  |  |  |  |
| 3. |  |  |  |  |
| 4. |  |  |  |  |
| 5. |  |  |  |  |

5.  Legally Binding Declaration.

I hereby declare that the information supplied on this form is complete and accurate.  The despatch will be in accordance with

the appropriate regulations covering Packaging, Transportation and Labeling of Dangerous Substances.

Name of Organisation: __________________________________________________________________________________

Address: _____________________________________________________________________________________________

_______________________________________		Post Code:           _______________________________

Tel.:        _______________________________________		Fax.:		 _______________________________

Name:    _______________________________________		Job Title:              ________________________________

Date:      _______________________________________		Company Stamp:  _______________________________

Legally Binding Signature: ________________________

# Text boxes

Fuse 1

Gas  Supply

(High Pressure

Cylinder)

Exhaust

(Chamber Out)

Power In

Vent

(cool Out)

Bleed

(flow Gauge)

Fine Bleed

(black)

Bleed

(red)

Chamber Out

(Black)

Chamber Fill

(Green)

Cool

(Blue)

Fill Level