**HealthyPhoton**

**TEC cooled HgCdTe (MCT)**

**Amplified Photodetector**

|  |
| --- |
| HPPD_M_B_wLOGO |


|  |  |
| --- | --- |
| **Date** | **Note** |
| 2019/5 | V1.0 |
|  |  |

AMPLIFIED MCT PD TEST REPORT

|  |  |  |  |
| --- | --- | --- | --- |
| **DETECTION MODULE** | | | |
| Preamplifier Model | TE0810-B-DC-20K | | |
| Detector type | PV8-TE | | |
| Preamplifier Power Supply | +5V, 3mA | | |
| **TEST CONDITIONS AND RESULTS** | | | |
| Parameter | | Unit | Value |
| Trans-impedance @Rload=1MΩ | | V/A | 20k |
| Low Cut-off Frequency | | Hz | 0 |
| High Cut-off Frequency | | KHZ | 150 |
| Output Noise Densivity @200kHz | | nV/ | 160.2 |
| Voltage Responsivity ±20%( 8um) | | V/W | 5.2E+2 |
| Detectivity ±20%(8um)@200kHz | | cm/W | 6.1E+8 |
| Dector temperature | | K | 223 |


**DETECTOR HEAD TEST REPORT**

|  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| **DETECTOR DESCRIPTION** | | | | | | | | | | | | |
| Active Area | | 1mm ×1mm | | Cooler Current | | 1.00A | | | Ambient Temperature | | 293 K | |
| Window | | wZnSeAR | | Cooler Current Max | | 1.20A | | | Test Date | | 10.6.2019 | |
| **TEST CONDITIONS AND RESULTS** | | | | | | | | | | | | |
| **Detector**  **No** | **Reverse Bias Voltage** | | **Detector Temperature** | | **Thermistor Resistance** | | **Detector Dynamic Resistance** | **Current Noise Density** | | **Responsivity ±20%( 8um)** | | **Detectivity**  **±20%( 8um)** |
| **mV** | | **K** | | **KΩ** | | **Ω** | **pA/Hz1/2** | | **A/W** | | **cmHz1/2/W** |
| **90621** | 0 | | 223 | | 91 | | 669 | 8.051 | | 0.026 | | 6.1E+8 |



|  |
| --- |
| **SPECTRAL/ELECTRICAL CHARACTERISTICS** |
| **IMG_5970_副本** |

|  |
| --- |
| **To package and Lead** |
|  |



**Order information**

**HPPD- XX XX XX - X - XX - XXX**

**① ② ③ ④ ⑤ ⑥**

**Note:**

**①=TE detector with TEC cooler**

**①=RT detector without TEC cooler**

**②=03, 04, ...or 12 detector with cut-off wavelength at 3, 4, ...or 12um**

**③=02,05,10,20 detector active area is 0.2, 0.5, 1.0, 2.0 mm2**

**④=A A type preamplifier seperate TEC driver**

**④=B B type preamplifier with integrated TEC driver**

**⑤=DC DC coupled preamplifier**

**⑤=AC AC coupled preamplifier**

**⑥=01K, 02K, ...or 20K Preamplifier gain is 1k, 2k, ... or 20k**

**For example:**

**HPPD-TE0810-B-DC-20K means the customer would like to order an 1mm2 active area MCT detector with TEC cooler at cuf-off wavelength of 8um, with a B type DC coupled preamplifier, the preamplifier gain is 20K.**

**Cautions:**

**1: The above product specifications are subject to change without notice.**

**2: The suitable ESD protecting measures are recommend in storage,transporting and using.**