| 品牌： | ROOKO瑞柯 | 税 率 |
| --- | --- | --- |
| 请选择对应 销售公司 | 宁波瑞品仪器有限公司 | 开具13%增值税专用发票 |
|  | 宁波盘羊仪器有限公司 | 开具3%增值税普通发票 |
| 开票品名/规格： | LX-9831电压降测试仪 |

LX-9831电压降测试系统

LX-9831 Voltage drop test system

- 概述Overview：

电压降、接触电阻、温升是车用电线束插接器在强电流电路和弱电流电路下电性能评定指标；集成电路芯片控制，恒流源供给输出，液晶显示，提供USB、232通讯接口；可外接控制单元；选购PC软件系统，分析电流、时间、温升、接触电阻和电压降变化曲线及数据模型；适用于科研院所、企业研发、进料检验和质量部门.

Voltage drop. contact resistance and temperature rise are the electrical performance evaluation indexes in the circuit of strong current circuit and weak current circuit.Integrated circuit and high-precision chip control. constant current supply output. LCD display. Provide USB. 232 communication interface.External control unit.Choose to buy PC software system. analyze current. time. temperature rise. contact resistance and voltage drop curve and data model. It is suitable for providing more accurate and reliable data assurance for technical departments such as research institute. enterprise R&D.

二.原理Principle：

通过电流、电压和时间，来检测电线束插接器电压降、接触电阻和温升变化关系，是汽车用插接器的电气性能指标.

Based on the current. voltage and time to detect Voltage drop of Electrical beam splitter .Relation between contact resistance and temperature rise change.

三.适用范围Scope of application：

用于电线束、插接器、开关触点、继电器、汽车开关等相关产品之接触电阻、温升、电压降或负载测量，同时也适用于具有类似特性产品使用.

Widely used in measuring contact resistance. temperature rise. voltage drop. or load of  electric wire beams. plug. switch contact. relays. automotive switch and other related products. Also is suitable for use with similar characteristics products.

四.参照标准Standards：

：QC/T 730 ;ISO 6722：2002  GB/T 2951; GB/T 3048； QC/T29106; QC/T413 ;QC/T413-2002；SAE/USCAR38 2009 .

五.参数资料： data

| 测量项目/ Measurement items | 规格/ Technical specifications | 分辨率 Accuracy |
| --- | --- | --- |
| 1.测试电流DC（选购 Test current(choose to buy) | 0-60A | 0-100A | 0-200A | -1000A | 0.01A |
| 2.测试电压（开路） test voltage(open circuit) | DC0-30V | 0.01V |
| 3.电压降范围voltage drop range | 0-1999.99mV | 0.01mV |
| 4. 测试时间test time | 测试时间0.01S—99.99H Test time 0.01S—99.99H |
| 5．温升范围temperature rise range | 常温-500度，精度0.5度 Normal temperature -500 degree，accuracy 0.5 degree. |
| 6.接触电阻contact resistance | 0.1μΩ—2KΩ.分辨率resolution 0.01μΩ |
| 7. 统计功能Statistical function | 统计计数：0-99999 Statistical count for good and bad products:0-99999 |
| 8. 测量模式Measurement mode | 一般材料类和电线束类General materials and wire harness |
| 7. 电压降设定 Voltage drop setting | 上下限设定功能及声光报警 Upper and lower limit setting function and |
| 8.设备管理Device Management | 仪器管理密码设定，自我校准，数据管理。存储功能 Instrument Manager password settings. self-calibration. data management.storage function. |
| 校准分流器（选购） Calibration shunt(choose to buy) | 0.2级标准分流器75mv自我校准仪器，规格如下：1）60A.75mv；2）100A.75mv；3）200A.75mv；4）其他规格 0.2 standard shunt 75mv self-calibration instrument. specifications are as follows: |
| 10.语言language | 提供中文或英文两种供选择providing Chinese or English language versions available. |
| 11.电源输入power supply | AC220V+10%;380V |
| 12.环境要求 Environmental requirements | 仪器使用环境要求Environmental requirements：工作温度：10℃－30℃，湿度：≤65%RH |
| 13. 治具fixture | 依据产品选购治具1付或者定制Choose to buy fixture 1 pay based on product or customized. |
| 14.外部单元External unit | 提供USB;232接口；选购：PC软件，电脑，打印机及其他控制单元Provide USB. 232 interface. choose to buy: PC software. computers. printers and other control units. |

配套方案：解决各材料状态 --固态、液态、气态、颗粒状 电阻、电阻率、电导率测量

ROOKO瑞柯仪器-------专注于新材料测量与分析仪器解决方案

----解决材料：

导体、半导体、绝缘材料常温及高温等环境下电性能分析.

----我们一直在做：

研发、生产、销售研发、生产、销售、租赁、实验室样品分析及后延扩展服务.