产品名称：高低温恒温水浴

型号：HWY-30

一、产品简介

本水浴适用于各种试件的养护，控温范围在常温条件以下使用的仪器，具有制冷速度快、噪音低及恒温精度高等优点。用途广泛，可与针入度连接做符合各种温度试验。是检测和试验单位的设备。

1. Product introduction This water bath is suitable for the maintenance of various specimens. The temperature control range of the instrument used under normal temperature conditions has the advantages of fast cooling speed, low noise and high temperature accuracy. It has a wide range of uses and can be connected with the penetration degree to meet various temperature tests. It is the equipment of testing and testing unit.

二、技术参数

1、设备控温范围:3-60℃

2、电源电压220v士10v

3、显示精度：±0.1℃

4、水浴尺寸:460X320X300

5、加热功率:2000w

6、制冷功率:280W

7、总重：63kg

8、外形尺寸长660mm宽650mm高1050mm

2. Technical parameters 1. Equipment temperature control range: 3-60℃ 2. Power supply voltage 220v ± 10v 3. Display accuracy: ±0.1℃ 4. Water bath size: 460X320X300 5. Heating power: 2000w 6. Cooling power: 280W 7. Total weight: 63kg 8. Dimensions: length 660mm, width 650mm, height 1050mm

三、安装说明：

1、设备要安放置在室温10℃到30℃；相对湿度不应高于80%的室内。室内应有良好的通风。

2、拆箱时要注意不要磕碰箱体及部件。

3、本设备电源建议用带漏电保护的空气开关，

4、本设备应放量在平整的地面，并且要求设备底部的四个足要全部接触地面并且着力均匀。

3. Installation instructions: 1. The equipment should be placed in a room with a room temperature of 10℃ to 30℃; the relative humidity should not be higher than 80%. The room should have good ventilation. 2. Be careful not to bump the box and parts when unpacking. 3. It is recommended to use an air switch with leakage protection for the power supply of this equipment. 4. The equipment should be placed on a flat ground, and all four feet at the bottom of the equipment should be in contact with the ground and the force should be even.

四、调试说明：

1、接通电源后将水箱内加上水，对设备进行调试，加水水位不能低于水箱内的篦子。（注意：不加水不能开机！）

2、加满水后请检查设备各个部位是否有漏水现象。如果有漏水现象请紧固该部位

螺栓或通知本公司

3、开机，依次按下电源键=》制冷键=》水泵键。这时水泵搅拌装置启动。但是，压

缩机不会启动(说明：压缩机在出厂时已经设量了延时保护装置，一般需要5-8分钟才

启动)

4、试机温度为：依次为25℃、15℃、5℃。首先从25℃开始，当温控表显示温度到达

25℃后，加热系统会自动运行，而制冷系统会始终运转，靠加热系统来补偿由于制冷系

统带来的热量损失，（我公司技术部之所以这样设计线路原因：其一，为了更精准的控制

温度，其二，为了保护压缩机并延长压缩机的使用年限)。这时水浴内水的温度还会继较

下降到大概低于设置温度3到5℃，然后逐渐上升直至到达所设置的控温温度并恒温。

15℃与5℃的试机过程依次类推。

试机完毕后将水浴内的水放出，并把各个部位擦拭干净以备后用。

Four, debugging instructions: 1. After the power is turned on, add water to the water tank to debug the equipment. The water level cannot be lower than the grate in the water tank. (Note: You can't boot without adding water!) 2. After filling up with water, please check whether there is water leakage in all parts of the equipment. If there is water leakage, please tighten the part Bolt or notify the company 3. Turn on the machine and press the power button = "cooling button =" water pump button in turn. At this time, the pump stirring device is activated. However, pressure The shrink machine will not start (Note: The compressor has a delay protection device set at the factory, and it usually takes 5-8 minutes to start up) 4. The test machine temperature is: 25°C, 15°C, and 5°C in sequence. First start at 25℃, when the temperature control meter shows that the temperature reaches After 25℃, the heating system will operate automatically, while the refrigeration system will always operate. The heating system is used to compensate the refrigeration system. The heat loss caused by the system, (the reason why the technical department of our company designed the circuit in this way: First, for more precise control Temperature, secondly, in order to protect the compressor and extend the service life of the compressor). At this time, the temperature of the water in the water bath will continue to be higher Decrease to about 3 to 5°C below the set temperature, and then gradually increase until it reaches the set temperature control temperature and keeps constant. The test machine process at 15°C and 5°C is analogized in turn. After the test machine is finished, let out the water in the water bath and wipe all parts clean for later use.

volume_up

content_copy

share

五、操作步骤：

1、将预制好的试样整齐的排列于工作室内部的托板上，向工作室内加满水，使水

浸没试件并高出试件30mm左右。然后接通电源依次打开制冷搅和水泵开关。

2、将所需温度设定在温控仪表上。设定方法如下：按下SET键随即松开会看到下

排设定显示部位的数码管出现闪动状态。按△增加，按V设定值减小，设定亮毕后，

按SET键返回。温控表将进入自动控制程序。

3、当温度达到到控温仪的所需控制温度时，温控仪会自动恒温。此时红灯亮并

使水箱温度稳定在该温度上，此时记下时间。

4、试件在水箱内温度恒定至少30分钟后方可进行试验。

5、试验结束后，旋动水箱底部的放水阀将箱内水放净，擦净仪器表面水份将仪器

放置干燥通风处。

5. Operation steps: 1. Arrange the prefabricated samples neatly on the pallet inside the studio, and fill the studio with water to make the water The specimen is immersed and is about 30mm higher than the specimen. Then turn on the power supply and turn on the cooling stir and water pump switch in turn. 2. Set the required temperature on the temperature control instrument. The setting method is as follows: press the SET button and then loosen the meeting to see the next The digital tube on the display position of the row setting flashes. Press △ to increase, press V to decrease the setting value, after the setting is bright, Press SET to return. The temperature control meter will enter the automatic control program. 3. When the temperature reaches the required control temperature of the temperature controller, the temperature controller will automatically maintain a constant temperature. The red light is on and Make the temperature of the water tank stabilize at this temperature, and write down the time at this time. 4. The test piece can be tested only after the temperature in the water tank is constant for at least 30 minutes. 5. After the test is over, turn the drain valve at the bottom of the water tank to drain the water, and wipe the surface of the instrument to remove the water Place in a dry and ventilated place.

volume_up

content_copy

share

六、注意事项：

1、本产品使用时应该有良好的接地。

2、水浴后盖距离墙面不能少于50cm。

3、水槽内应该加人含矿物质很少的饮用水，以免水箱及循环管件内结垢。

4、使用电压应该严格控在220V±10V。

5、使用中不可以开关缩机，使用亮毕后依次关闭水泵、制冷、电源开关。

6、使用完毕后将水浴内的水放出，并把各个部位擦拭千净以各后用。

7、压缩机在长期不使用的情况下很容易出现故障，所以本设备应该每个月进行一

次试运行，以防压缩机损坏。

6. Matters needing attention: 1. This product should be well grounded when using it. 2. The distance between the back cover of the water bath and the wall shall not be less than 50cm. 3. Drinking water with few minerals should be added to the water tank to avoid scaling in the water tank and circulating pipe fittings. 4. The operating voltage should be strictly controlled at 220V±10V. 5. Do not turn on or off the shrink machine during use, turn off the water pump, refrigeration, and power switches in turn after use. 6. After use, release the water in the water bath, and wipe all parts for later use. 7. The compressor is prone to failure when it is not used for a long time, so this equipment should be carried out once a month. Test run again to prevent damage to the compressor.