SKALAR

自动分析过程中的最佳伙伴

| 荷兰skalar Primacs SN 总氮-蛋白质分析仪 快速、精确、无需有害化学试剂、全自动分析 |  |
| --- | --- |

| Primacs SN总氮-蛋白质分析仪 |
| --- |
|  | 应用于: 乳制品 谷类 肉制品 食品 植物&土壤 肥料 烟草 啤酒&麦芽 动物饲料 制药 等等 |
|  | 对于食品、谷物、乳制品、肥料、动物饲料、植物、烟草及相关产品，全氮和蛋白质含量的测定是非常重要的。传统凯氏定氮法作为工业标准已执行很多年，但繁复的操作步骤、耗时长、使用腐蚀性酸、难以把有机物定量转变成氨以及严重的环境污染较难满足当前检测发展的需求。在过去的十年多里，一种快速、精确、低成本、无污染、自动化高的定氮方法很快的在欧美国家得到了广泛应用，成为凯氏法的替代方法， 即杜马斯（DUMAS）燃烧定氮法。杜马斯法目前已被很多组织认可，如: AOAC,ASBC, EBC, AACC和 ISO，成为法定的氮/ 蛋白质分析方法。在美国、加拿大和德国的某些领域,杜马斯法甚至作为唯一的定氮标准。 杜马斯法与传统凯氏法的定氮方法比较 传统凯氏法 杜马斯法 所需的化学试剂 浓H2SO4, 40%NaOH, K2SO4, CuSO4,乙醇,甲基红,溴甲酚绿, 硼酸，蒸馏水 空气,氧气,氦气,氧化铜, 含氮标准物质,氮催化剂, 干燥剂,氧化铝 所需的仪器设备 凯氏定氮仪,凯氏烧瓶,消煮炉, 分析天平,通风橱 杜马斯定氮仪,气瓶,分析天平 化学废液的处理 必须收集并经过专业人员处理 无化学废液 测试时间 120min/炉，手动蒸馏、滴定 3-5min/个,全自动连续进样分析 自动化程度 消化、蒸馏、滴定，手工分析 全自动分析 危险程度* 6 2 精确度（CV%） 1.2 0.7 注：* 危险程度划分为10 个等级（1-10），10 为极端危险，1 为完全安全。 荷兰skalar公司具有30多年用于蛋白质分析的定氮法自动化的检测仪器的制造经验，其Primacs SN全自动燃烧法蛋白质-总氮分析仪采用杜马斯燃烧法，整个测试过程仅需几分钟时间，专利的垂直式燃烧炉和自动进样器垂直进样技术减少外界干扰及日常维护的需要，具有分析速度快、结果精确、全自动化操作、无危害和无污染分析的特点，广泛应用于土壤、植物、食品、动物饲料、制药等领域 |

操作原理

分析过程：

- 非均质样品在称样之前必须通过样品磨粉碎过筛处理。

- 样品称重后数据自动装载至软件中，消除人为录入的误差和计算误差。

- 样品与纯氧在1100ºC下燃烧，，燃烧生成的气体被载气携带直接通过氧化铜（作为催化剂）而被完全氧化。此外，化合物中一定量的难氧化部分会被载气携带通过催化剂反应进一步氧化，产生的气流通过高效无需维护的Peltier电子制冷器除去所有水分。

- 在精确稀释后，该气流收集在收集容器中，并通过还原炉将氧化氮还原成氮。

- 气流通过捕集器除去卤素等杂质后，氮的含量通过灵敏精确的热导检测器检测。

- Skalar的专用软件包全程控制仪器、计算数据、输出Q.C报告、数据可很简单地传输到LIMS系统中。

技术指标

| 分析项目 | 总氮-蛋白质 |
| --- | --- |
| 方法 | 根据DUMAS方法，高温催化燃烧，热导检测器检测（TCD） |
| 样品 | 乳制品、谷类、食品/饮料、植物、土壤、肥料、烟草、啤酒/麦芽、动物饲料、制药等 |
| 测量范围 | 0 - 200 mg N绝对值 |
| 检测限 | < 0.01% (< 100 ppm) |
| 分析时间 | 约 3 - 5 分钟 |
| 样品重量 | 最大1g（通常100-300mg）或液体样品体积1 ml |
| 精密度 | < 0.5 % RSD（大约300 mg 被测物质）在10 % N含量时 |
| 数据处理 | 面积计算（多点线性回归），天平接口数据自动录入，原始数据储存和可连接至LIMS |
| 自动取样器 | 20样品位自动取样器 |
| 气体 | 载气： He(99.99%);燃烧气：O2(99.995%) |
| 反应炉温度 | 燃烧炉：1100℃； 氧化炉：750℃； 还原炉：600℃ |
| 符合方法 | AOAC 990.03, AOAC 993.13, AOAC 968.06, AOAC 992.15, AOAC ，992.23, ASBC, AACC 46-30, FGIS, AOAC, CGC， DIN / EN / ISO 14891 ，DIN/ISO 13878, LUFA, MEBAK, ICC 167，ISO/FDIS 16634 ，AOCS Ba 11-65 |
| 电源 | 220 - 240 V, 50/60 Hz (最大 2000 VA) |
| 尺寸(hxdxw) | 83 x 50 x 59 cm (32.7 x 19.7 x 23.2 英寸) |
| 重量 | 56 kg |

Primacs SN 分析仪构造

电子制冷器

捕集器       									                       灵敏的TCD检测器

| Primacs SN 软件 |
| --- |
|  | Primacs SN分析仪和专用的SNAccess软件采用Windows平台便于使用和确保兼容LIMS。容易使用的软件具有多种功能便于各人快速和简单地开启任何分析。 此外，该软件显示和储存实时信息如温度和检测器信号，进行可靠的元素分析。当无人监控操作时用户可设定警报限值安全操作。 可由用户对分析仪设定选择不同的样品类型，分析仪校正和样品分析格式可简单设置和储存。可通过接口将天平的数据自动录入到软件中，分析时可添加样品、删除或编辑样品位置。 最终的数据可和容易地输出和打印，所有可提供数据、单点或多点的校正曲线图形数据可按客户需要编制输出格式、显示及打印。所有的分析数据均可以文本形式输出，便于外部程序评估。 |

| 蛋白质/总氮典型的检测结果 |
| --- |
| 样品类型 小麦面 奶制品 麸质 小麦庄稼 全脂黄豆 鱼粉 大豆油饼 样品类型 草 植物 茶叶 黄瓜叶 番茄叶 树皮 香蕉叶 混合肥料 草莓叶 莴苣叶 土壤 | 样品重量 100 – 200 mg 100 – 200 mg 100 – 200 mg 100 – 200 mg 100 – 200 mg 100 – 200 mg 100 – 200 mg 样品重量 100 – 200 mg 100 – 200 mg 150 – 200 mg 100 – 200 mg 100 – 200 mg 100 – 150 mg 100 – 150 mg 50– 100 mg 100 – 200 mg 100 – 200 mg 200 – 300 mg | 重复分析次数 7 5 5 7 5 5 5 重复分析次数 5 5 5 5 5 5 5 5 5 5 10 | 蛋白质 % 10.69 15.00 23.75 9.56 41.75 51.94 50.12 总氮% 4.31 1.74 2.54 5.23 4.50 0.78 3.40 3.52 2.26 4.55 0.18 | RSD % 0.16 0.90 0.41 0.94 0.78 1.09 0.79 RSD % 0.23 0.65 0.79 0.72 0.52 0.98 0.07 0.71 0.87 0.53 1.05 |

典型应用

鲜奶中的总氮/蛋白质分析 燃烧炉温度：1100ºC/ 氧化温度：750ºC/ 还原温度：600ºC / 分析时间：310 秒 / 校正物质：EDTA 样品制备：无需制备，鲜奶放入石英坩埚中直接燃烧和检测 蛋白质换算因子：6.25 杜马斯 Vs 凯氏定氮（r=0.994）

动物饲料中的总氮/蛋白质分析 燃烧炉温度：1100ºC/ 氧化温度：750ºC/ 还原温度：600ºC / 分析时间：310 秒 / 校正物质：EDTA 样品制备：对于湿的饲料，在粉碎前先风干混匀，样品粉碎过1mm筛，注意在粉碎期间防止样品过热 蛋白质换算因子：6.25 杜马斯 Vs 凯氏定氮（r=0.993）

麦芽中的总氮/蛋白质分析 燃烧炉温度：1100ºC/ 氧化温度：750ºC/ 还原温度：600ºC / 分析时间：310 秒 / 校正物质：EDTA 样品制备：称50g麦芽粉碎后过600µm筛(Sieve No. 30)，水份需进行修正 蛋白质换算因子：6.25 杜马斯 Vs 凯氏定氮（r=0.9993）

典型应用

烟草样品中的总氮分析 燃烧炉温度：1100ºC/ 氧化温度：750ºC/ 还原温度：600ºC / 分析时间：310 秒 / 校正物质：EDTA 样品制备：样品在60℃下风干除水，样品粉碎后过1.0mm筛，注意在粉碎期间防止样品过热 杜马斯 Vs 凯氏定氮（r=0.992）

植物的总氮/蛋白质分析 燃烧炉温度：1100ºC/ 氧化温度：750ºC/ 还原温度：600ºC / 分析时间：310 秒 / 校正物质：EDTA 样品制备：样品在60℃下风干除水，样品粉碎过0.5mm筛，注意在粉碎期间防止样品过热 蛋白质换算因子：6.25 杜马斯 Vs 凯氏定氮（r=0.995）

土壤的总氮析 燃烧炉温度：1100ºC/ 氧化温度：750ºC/ 还原温度：600ºC / 分析时间：310 秒 / 校正物质：EDTA 样品制备：在40 º C下风干,混匀均质化样品,粉碎过2.0 mm筛 杜马斯 Vs 凯氏定氮（r=0.998）

| PrimacsATC总碳分析仪 |
| --- |
|  | 应用于: 食品 植物 土壤 肥料 淤泥 沉积物 等等 |

在食品、土壤&植物、沉积物和淤泥等样品中分析碳是很重要的。skalar PrimacsATC总碳分析仪采用高温催化燃烧法，可快速和精确地分析固体、液体样品中的总碳。

样品在可重复使用的石英坩埚中称重，并装载在内置的自动取样器中，重量值可通过天平与仪器联机自动传输至软件的工作表格中，避免产生错误。专利的“bottom-to-top”垂直进样技术将样品提升至高温燃烧炉中,在1050°C下催化燃烧，所有的碳转化成CO2, CO2的测量采用非分散红外检测器(NDIR)后得出总碳含量。

软件可同时显示出碳的实时峰，结果可很方面地打印和输出到LIMS系统中，在运行期间，可随时插入新的分析样品。灰化后的样品保留在坩埚中并在分析后自动无污染地转移到仪器外部，这样避免了样品灰化后残留在燃烧区而因此减少了维护。高温燃烧炉组合灵敏的双范围红外检测使得PrimacsATC可分析由ppb级至百分含量水平的碳含量的分析

技术指标

| 分析项目 | 总碳 |
| --- | --- |
| 分析方法 | 高温催化燃烧法 |
| 检测方法 | 非分散红外检测器（NDIR） |
| 应用 | 土壤、植物、食品、液体和过氧化氢样品等 |
| 自动进样器 | 20位随机取样器，垂直进样 |
| 分析范围 | 0.01 – 120 mg C绝对值 |
| 测量范围 | 0.001 – 100% |
| 分析时间 | 3-5分钟 |
| 样品尺寸 | 固体样品：最大1g(通常100-300mg)； 液体样品：最大200 mg (μl) |
| 重复性 | < 1 % RSD |
| 数据处理 | 面积计算（多点线形回归），天平称重界面，原数据储存和连接到LIMS |
| 所需气体 | 燃气：99.95 % 氧气, 300 kPa. |
| 反应炉温度 | 燃烧炉:1000℃; 催化炉：750℃ |
| 电源 | 220V – 240 V/50/60Hz |
| 功率 | 1500 VA |
| 尺寸(hxdxw) | 83 x 50 x 59 cm, (32.7 x 19.7 x 23.2 inches) |
| 重量 | 56 kg |

###### TOC add-on module for Solid samples

The PrimacsSERIES Carbon / Nitrogen analyzers have now been extended with the new PrimacsMCS add-on module for Total Organic Carbon (TOC) analysis for solid materials such as soil, sediments and sludges. The PrimacsMCS operates in combination with the FormacsSERIES Total Organic Carbon analyzer (liquid samples), offering an economical concept to environmental laboratories for Total Organic Carbon analysis on aqueous samples as well as solid materials.

The PrimacsMCS is equipped with a dual-oven design, allowing separate analysis of Total Carbon (TC) and Inoragnic Carbon (IC). Total Carbon is determined by catalytic oxidation of the sample at 1100Â°C, converting the carbon present in the sample to CO2, which is detected by the NDIR detector of the FormacsSERIES Total Organic Carbon analyzer. Inorganic Carbon is determined by acidification of the sample in the IC reactor, which converts the Inorganic Carbon present to CO2. The Windows based PrimacsMCS software collects the data and calculates the TOC concentration of the sample by substraction ( TC - IC = TOC). The PrimacsMCS analyzes TC, IC and TOC in a range of 100 Âµg to 40 mg absolute carbon in up to 3 g of sample weight. Due to the automatic low-to-high range switching IR detector of the FormacsSERIES Total Organic Carbon analyzer, the combination of the FormacsSERIES analyzer with the PrimacsMCS add-on module offers environmental laboratories a powerful tool to analyze TOC from low ppb concentrations in aqueous samples up to % levels in solid materials. The PrimacsMCS conforms to international regulations such as EPA, CEN, ISO, NEN-EN 13137 and US EPA 9060

#### PrimacsSLC

###### Total Organic Carbon Analyzer, for Solid & Liquid Samples

The increasing demand for reliable TOC analysis in solid materials such as sludges, sediments and soils have led Skalar to complement its range of Total Organic Carbon analyzers with the PrimacsSLC Analyzer. Based on a dual oven design the analyzer is capable of performing rapid, reliable, separate determinations of Total Carbon (TC), Inorganic Carbon (IC) and Total Organic Carbon (TOC) on one instrument, and has been developed especially for solid materials on sample weights of up to two grams. Automatic balance interfacing avoids the necessity of precise weighing and transfer of weight data. The reusable quartz crucibles combined with the unique vertical sample introduction system and fully computerized operation/calculation, make the PrimacsSLC the most economical choice in today"s environmental laboratory.

The PrimacsSLC conforms to international regulations such as CEN, ISO and NEN-EN 13137 and US EPA 9060

#### PrimacsSN

###### Protein - Nitrogen Analyzer

Skalar's New Protein - Nitrogen Analyzer is the modern alternative for nitrogen analysis in the busy routine laboratory. The combustion (Dumas) method used in this new Skalar analyzer is rapidly becoming the accepted alternative for the old fashioned Kjeldahl method. The method is fast, safe, economical and easy to operate. 
The PrimacsSN is applicable in a wide variety of application areas including soil and plant, food, brewing industries, and environmental waste, to name a few.

The PrimacsSN conforms to international regulations such as ISO 15670, ISO 13878, EN 61010, DIN 10467, AOAC 990.03 , AOAC 992.15, AACC 46-30, AOCS Ba 11-65

The sample is automatically introduced from a 20 position sample tray into the unique vertical positioned combustion oven , the resulting gas mixture passes through a second oxidation oven whereby all nitrogen compounds are converted into NOx. After this the Peltier cooler removes all water before passing the copper reduction column removing excess O2 from the gas flow, and converting NOx to N2. The stream passes through a CO2 scrubber and subsequently through a magnesium perchlorate H2O scrubber. The resulting gas stream, consisting of N2 and He (carrier gas), passes the thermal conductivity detector which determines the nitrogen content. The Analyzer is fully computer controlled with result calculation, QC and flexible reporting.

Protein Flow Diagram (Flash animation)

#### PrimacsSNC

###### Total Carbon and Total Nitrogen / Protein Analyzer for solid and liquid samples

The analysis of Carbon and Nitrogen is important in soil, plant, animal feed, food samples, sediments and sludge samples. The Skalar PrimacsSNC Analyzer combines the analysis of Total Carbon (TC) and Total Nitrogen (TN) in one unit using high temperature catalytic combustion. The samples are introduced into the high temperature combustion oven by the unique vertical "bottom-to-top"sample introduction system.

At 1050Â°C the carbon is completely oxidized to CO2 in the presence of a catalyst. The CO2 is measured by Non Dispersive Infra Red Detection (NDIR) for Total Carbon. The analysis of Nitrogen is based upon the well-proven DUMAS technology. Nitrogen is converted in NxOy which is is reduced at 600Â°C to N2. The N2 gas is measured by Thermal Conductivity Detection (TCD). The PrimacsSNC analyzer provides an accurate and reliable solution for the automation of Total Carbon and Total Nitrogen analysis and has been designed as an easy-to-use and low maintenance analyzer. Due to the unique vertical "bottom-to-top" sample introduction the sample ashes remain in the crucible after the analysis and are taken out of the instrument with removal of the crucible. This avoids sample ash build-up in the combustion zone and therefore reduces the maintenance. The re-usable crucibles and the low consumption of combustion and carrier gas result in a low cost-per-test.

#### Typical Applications

| PRIMACSSNC总碳和总氮分析仪 |
| --- |
|  | 典型的应用 • 土壤 & 植物 • 淤泥和沉积物 • 食物 • 废物焚化 • 废水实验室 等等 |

在土壤、植物、动物饲料、食品样品、沉积物和淤泥样品中分析碳和氮是很重要的。skalar PrimacsSNC分析仪采用高温催化燃烧法，可分析固体、液体样品中的总碳和总氮。

分析原理

置于全自动进样器的样品采用专利的“bottom-to-top”垂直进样技术进入高温燃烧炉中,在1050°C下催化燃烧，所有的碳转化成CO2, CO2的测量采用非分散红外检测器(NDIR)后得出总碳含量.总氮的分析基于杜马斯(DUMAS)燃烧法,所有的氮转化成NxOy,在600°C催化还原成N2, N2的测量采用热导检测器(TCD)，软件可同时显示出碳和氮的实时峰，结果可很方面地打印和输出到LIMS系统中。在运行期间，可随时插入新的分析样品。样品在石英坩埚中称重，其重量值可通过天平与仪器联机自动传输至软件的工作表格中，避免产生错误。

PrimacsSNC分析仪提供精确可靠的全自动总氮和总碳分析，仪器容易操作，维护低。由于采用专利的垂直进样技术，灰化后的样品保留在坩埚中并在分析后自动无干扰地转移到仪器外部，这样避免了样品灰化后残留在燃烧区而因此减少了维护。可重复使用的坩埚和低消耗的燃烧装置和载气，分析成本低。

PrimacsSNC符合国际标准：CEN, ISO 10694, NEN-EN 131137, AOAC 990.03,AOAC 992.15, AACC 46-30, ASBC 等等

PRIMACSSNC技术指标

| 分析项目 | 总氮/蛋白质和总碳 |
| --- | --- |
| 总碳分析方法 | 高温催化燃烧和非分散红外检测器（NDIR）检测 |
| 总氮/蛋白质 | 杜马斯（Dumas）燃烧技术和热导检测器（TCD）检测 |
| 样品类型 | 固体或液体样品，食品、土壤、植物、淤泥、沉积物、废物等 |
| 自动进样器 | 20位自动进样器 |
| 进样 | 专利的“bottom-to-top”垂直进样系统 |
| 其它特点 | 自动天平数据录入接口 |
|  | 气体反吹系统可消除环境气体的干扰，保障低含量分析时的精度 |
|  | 可重复使用的石英坩埚 |
| 符合国际法规 | ISO, EN, AOAC, ASBC 等等 |
| 总碳的测量范围 | 0.1 – 400 mg 绝对值 |
| 总氮的测量范围 | 0.5– 100 mg N 绝对值 |
| 分析时间 | 3 – 5分钟 |
| 固体样品重量 | 最大1g（通常100-200 mg） |
| 液体样品 | 最大200 mg (µl) |
| 炉温 | 氧化炉：1050°C; 还原炉：680°C |
| 重现性 | < 1 % RSD, FSD（满量程） |
| 数据处理 | 面积计算 |
| 气体 | 燃烧气体：O2 (99.995%), 还原载气: He (99.99%) |
| 功率 | 220 – 240 V, 50/60 Hz |
| 尺寸(hxdxw) | 83 x 50 x 59 cm (32.7 x 19.7 x 23.2 inches) |
| 重量 | 58 kg |

| 广州办事处： 廣州市體育東路廣州南方證券大廈1206室 電話: (020) 38878521 / 38878522 傳真 : (020) 3887 9660 E-mail : gz@chelleson.com 北京办事处： 北京市朝阳区德外馬甸裕民路12号E-1座元辰鑫大厦802-804室 电话：(010) 82252310 / 82251162 传真：(010) 82251161 E-mail: cls_bj802@sina100.com 上海办事处： 上海市北京西路605弄57号嘉发大厦B栋9楼C座 电话：(021) 52289066 / 62172031 传真：(021) 62177274 E-mail: clssh@sh163.net | 荷兰SKALAR公司总部 The Netherlands Tel. +31 (0)76 5486 486 Fax. +31 (0)76 5486 400 Internet: www.skalar.com 中国总代理： www.chelleson.com 香港总部： 香港新界葵涌葵喜街38号都会坊12楼1210-1212室 电话：(00852)24136700 传真：(00852)24024858 E-mail: hk@chelleson.com | 东莞办事处： 广东东莞市长安霄边中惠阳光国际商城二栋25楼J室 电话：(0769) 5425061 / 5425030 传真：(0769) 5425092 E-mail: clsdg@163.com 云南办事处： 云南省昆明市滇池路365號藍山數碼1-1-701 电话：0871-4647029/6685815 传真：0871-4647029 E-mail:kmchelleson2003@sina.com 成都办事处： 成都市锦华路1号万达广场1单元1906室 电话：028- 84191657 / 84191658 传真：028- 84191659 E-mail:clscd@chelleson.com |
| --- | --- | --- |

# Text boxes

燃烧炉易于开启

精确的气体

分流装置

自动天平数据传输接口

样品前处理

分析控制和数据处理

20位自动进样器

600℃还原反应炉

1100℃燃烧炉

TCD检测器

捕集器

接受

容器

分流

稀释器

电子

制冷器

自动进样器

750℃催化炉